Spectral alignment: what the whole-corpus view showed

Checking all 9,623 pre-training spectra and the five evaluation cohorts — 15 September 2026

Your suggestion, implemented

You asked that we look at all 9,000+ spectra by eye rather than trust summary statistics. We did, but as a single image instead of a scroll: each spectrum is one pixel row, so the whole corpus and all five cohorts fit on one page (Figure 1, attached). A vertical stripe is a metabolite peak; where a stripe steps sideways, those spectra are misaligned. This is both feasible at 9,623 spectra and more sensitive than scrolling, because a 30-point shift is invisible in a single trace but obvious as a step across thousands of stacked rows.

Finding 1 — the corpus was aligned by data point, not by chemical shift

The left panel shows a clear horizontal break part-way down the corpus. The spectra had been put on a common point count rather than a common ppm axis, so studies acquired with different parameters were silently offset against each other. Anchoring every spectrum on its near-0 ppm feature fixed it: the interquartile range of displacement fell from 768 points to 35 — about one linewidth. The right panel is the result.

Finding 2 — the cohorts were not misaligned; they were on different axes

Shifting the cohorts onto the corpus never worked, and the reason was not the size of the shift. Their acquisition metadata, which we obtained this week, shows different spectral widths. A peak then occupies a different number of data points, and no shift can correct that — the spectra must be resampled.

| Dataset | Spec. width | Field | Correction applied | Agreement with corpus |
| --- | --- | --- | --- | --- |
| Pre-training corpus | 20.024 ppm | 600 MHz | 0 ppm re-referencing | — (reference) |
| Barth | 12.0308 ppm | 950 MHz | resampled x1.664 | -0.11 → +0.64 |
| BrC-T2D | 20.1587 ppm | 800 MHz | resampled x0.993 | +0.13 → +0.26 |
| MTBLS326 | 20.02 ppm | 800 MHz | none needed | +0.16 |
| MTBLS563 | 20.0 ppm | 700 MHz | none needed | +0.59 |
| TBI | unknown | unknown | none — excluded | n/a |

Barth was the extreme case: at 950 MHz over a 12 ppm sweep, its peaks are spread 1.66x wider in data points than the corpus's. Resampling turned the worst-agreeing cohort into one of the best.

What this changes, and what is still open

The corpus is now internally consistent, and four of five cohorts sit on its axis. This removes a confound that would have depressed every transfer result, so the planned experiments should be re-run on the corrected data. Still unresolved: the absolute chemical shift axis. The corpus carries no true reference standard — the feature near 0 ppm is a broad hump, not TSP — so each dataset is internally consistent but their absolute ppm values are not established. This does not affect comparisons within a dataset; it does block quantitative matching against an external reference library.

One methodological point worth flagging

Before the metadata arrived we tried four independent ways to recover the correct axis from the spectra alone. All four gave confident answers, and all four were wrong when checked against a case whose answer we later knew — one estimated Barth's scale factor as 0.98 where the truth is 0.60. The dense metabolite region makes a wrong alignment fit almost as well as the right one. The practical conclusion is that acquisition metadata is not optional for this project. We now have it for four of five cohorts; TBI has none, which is why it is excluded rather than corrected.